# Hacking area
We’d like to present here possible modification and of existing PSU design. Some of them can be done by simply adding missed parts that are marked as optional in the design, another will require replacing existing parts with a new ones. Additionally our (and your) ideas about further improvements will be also listed here.
- Hack #1 – Choose power input
- Hack #2 – Set output range
- Hack #3 – Disable low ripple mode of operation
- Hack #4 – Remove Arduino powering over USB
- Hack #5 – Isolate the USB port
- Hack #6 – Communication ports protection
- Hack #7 – Use external W5500 Ethernet module
## Hack #1 – Choose power input
The power board can be connected to the DC or AC power input (of course, only one at the time). For the EEZ H24005 a AC/DC power modules are chosen. Therefore the DC input is used. If you got only PCBs or assembled modules and not the completed device and have some spare mains transformer you can use it here. Construction requirements for transformer is really minimal – only one secondary winding per channel is needed. With right output voltage and current/power you can (re)use it for new bench power supply.
Mains transformer output has to be connected to the AC input and few additional parts are required that power board becomes functional. Here is what you need to do:
- Remove C1, C2, C3, L1, X1, ZD1 if they exists on the PCB
- Mount on bottom layer THT parts X2, C4, C5
- Mount on top layer SMT parts D1, D2, D3, D4, R1 and if you like ZD2 (reuse ZD1 if exists)
Caution: make sure that max. output voltage (Vmax) never exceed absolute maximum rating of both IC1 nor Q1. That will permanently damage them and value in question is 60 Vdc that is about 41.6 Vac. The power board is tested with mains transformer Talema 55186-P1S2 (nominal output is 40 Vac).
## Hack #2 – Set output range
The EEZ PSU is until now tested with two output ranges: 50 V, 3.12 A and 40 V, 5 A. The later is selected for the EEZ H24005 model. If someone what to use non-standard power input e.g. 36 V, 100 W LED driver power module or smaller transformer then one mentioned in #1 that will requires modifications of CV and/or CC control loop gain.
The gain of CV control loop can be changed by manipulating feedback loop resistor of IC6B. The max. output voltage will be reached when it’s output is equal to reference voltage that is +2.5 V in our case provided by precision voltage reference IC10. Therefore for our default setup of 40 V we need a gain (actually attenuation) of 2.5 / 40 = 0.625 what translated in fractional number gives us round ratio of 5/8. Now we have to choose two resistor’s value with that ratio and check if is correct with one of many online op-amp voltage and gain calculator that can be found here. One of main obstacle here is required resistor’s precision. You should go with 0.1% and low temperature coefficient for good results. Table that follows offers resistors names and value for two additional ranges together with default one of 40 V.

Range	Required gain	R46, R56	R47, R57
0 – 30 V	0.08333 (5/6)	75 kΩ	15 kΩ
0 – 40 V (default)	0.0625 (5/8)	100 kΩ	20 kΩ
0 – 50 V	0.05 (1/20)	130 kΩ	20 kΩ

Gain of CC control loop depends of IC8A gain but also current sensing resistor R66 that defines what voltage drop can be expected for the current that flow to the load. According to Ohm’s law it’s obvious that voltage drop (U = I * R) will be higher and we can expect a better precision for smaller currents. But that will also means higher power dissipation on that resistor. For example with 200 mΩ we’ll get 1 V for max. range of 5 A but dissipation in that case is 5 x 5 x 0.2 = 5 W that will be hard to cool down and therefore unacceptable. With selected 20 mΩ it is 500 mW but max. output is also ten time smaller 100 mV. Again, as in case of CV control loop gain mentioned 100 mV has to be amplified to 2.5 V for full scale. Therefore required gain is 2.5 / 0.1 = 25. Table that follows offers two additional ranges together with default one of 5 A.

Range	Required gain	R61, R69
0 – 3.12 A	40	3.9 kΩ
0 – 4.16 A	30	3.24 kΩ
0 – 5 A (default)	25	2.37 kΩ

The PSU firmware has to reflect changes on the Power boards, otherwise you can get unexpected output values. The recommended place for all user code changes is conf_user.h file.
For example, we would like to expand a little bit voltage range because we found that when channel is calibrated it cannot reach max. value, e.g. instead of 40 V, we got 39.95 V. That can be done by changing CV control loop gain (default is 0.625) by increasing R47, R57 to 22 kΩ to decrease gain to about 0.0615. In that case max. programmed value (for DAC 0xFFFF) will be 40.67 V. Firmware constants of interest for 40 V, 5 A configuration are CH_PARAMS_U_40V, CH_PARAMS_40V_5A and CHANNELS.

There is a two possible way how to configure firmware to take care of our changes. First is to simply redefine CH_PARAMS_40V_5A, by changing voltage range value from 40.0f to 40.6667f. Add the following code into conf_user.h:

<pre><code class="arduino">
#undef CH_PARAMS_U_40V
#define CH_PARAMS_U_40V 0.0f, 0.0f, 40.0f, 40.6667f, 0.01f, 0.1f, 5.0f, 0.15f, 19.1f, 38.0f, 25.0f, CH_PARAMS_OVP
</code></pre>

Another way is to introduce new voltage parameters constant (e.g. CH_PARAMS_U_40.67V) and new channel parameters (CH_PARAMS_40.67V_5A) that will be used to set new definition of both channels in CHANNELS:

<pre><code class="arduino">
#define CH_PARAMS_U_40.67V 0.0f, 0.0f, 40.0f, 40.6667f, 0.01f, 0.1f, 5.0f, 0.15f, 19.1f, 38.0f, 25.0f, CH_PARAMS_OVP
#define CH_PARAMS_40.67V_5A CH_PARAMS_U_40.67V, CH_PARAMS_I_5A, true, CH_PARAMS_OPP_DELAY, 0.0f, 155.0f, 155.0f, 48.0f, 2.2f, 25.0f, 160.0f

#undef CHANNELS
#define CHANNELS \
    CHANNEL(1, CH_BOARD_REVISION_R5B9_PARAMS, CH_PINS_1, CH_PARAMS_40.67V_5A), \
    CHANNEL(2, CH_BOARD_REVISION_R5B9_PARAMS, CH_PINS_2, CH_PARAMS_40.67V_5A) \
</code></pre>
## Hack #3 – Disable low ripple mode of operation
The Power board offers MCU managed low ripple mode of operation. If such feature is not required or you are going to experiment with firmware that could potentially introduce some disturbance in limiting output current and power (Q1 or Q4 could be damaged in that way) then you can permanently or temporarily fix that by disabling it on the Power board directly. All what is needed is the following:
- Remove Q3
- Mount R13
## Hack #4 – Remove Arduino powering over USB (Obsolete)
This hack is related to Arduino Shield r1B9 only

The Arduino shield carry Arduino board used for digital control of the PSU. The Arduino board is also powered with 5 V via Arduino shield provided by Auxiliary power module. But Arduino boards are designed to be powered also from connected host PC via USB port. Therefore if USB cable is connected to the PSU, the Arduino board will stay powered even if PSU is disconnected from the mains voltage. That makes SW1 on the front panel effectively unusable.
If such behavior is not acceptable it can be easily fixed by removing mosfet T2 (see Fig. 1). Additionally, the IC2 could be removed since Vin power input is not used.

Fig. 1: Excerpt from Arduino Due schematic
Positions of mentioned parts is red marked on Fig. 2. When hot-air is used for desoldering additional care should be taken to not melt down 36-pin 0.1" vertical socket (far right).

## Hack #5 – Isolate the USB port
Fig. 2: Arduino Due with removed parts
If previous hack is too radical there is another possibility that not only separate USB powering but also isolate USB data lines making PSU completely galvanically isolated from connected host PC. That can be accomplished by populating Auxiliary power module with the following parts:
- USB isolator (IC20)
- Decoupling capacitors C100, C101, C103, C104
- Resistors R120, R121
- Ferrite beads FB6, FB7 and
- 5-pin header X10 (or X9 can be moved on that position)
USB cable that connects Arduino board with X9 has to be moved to X10.
## Hack #6 – Communication ports protection
When previous hack is applied it reduce greatly a chance that due to faulty condition PSU jeopardize connected host PC and vice versa. But regardless of that if USB isolation is added or not USB lines can be additionally protected using ESD protector IC. Same type of protection can be added to protect Ethernet lines connected to the IC29. Mentioned protection can be achieved by populating the following places on the Auxiliary power module:
- IC19, C102 for USB and
- IC18 for Ethernet
## Hack #7 – Use external W5500 Ethernet module (Obsolete)
The default mean of network communication with the PSU is by using W5500 Ethernet controller (IC29) located on the Arduino shield. If you are in process of assembling the PSU and have widely available Ethernet module as one shown on Fig. 3 you can use it temporarily for network connectivity testing even before IC29 and related parts are soldered. For connecting this module you’ll require the following:
- Solder 10-pin, 2x5 0.1” header (X20) and
- 10-pin IDC cable
If such module is planned to be used permanently take into account that existing EEZ H24005 model rear panel cannot be used for mounting this module directly because it’s presumed that RJ-45 connector is mounted on the Auxiliary power module. The X20 pinout is shown in the table below. Please note that it shares the same chip select as IC29 therefore they cannot be installed concurrently.
Fig. 3: W5500 Ethernet module

#	Pin name	Description
1	+3.3 V	Not used
2	+5 V	Ethernet module power input
3	MISO	SPI MISO signal
4	Gnd	Ethernet module power return
5	MOSI	SPI MOSI
6	RESET	Reset
7	ETH_SELECT	Chip select
8	ETH_IRQ	Interrupt request
9	SCLK	SPI clock
10	n.c.	Not used