# Technology overview

The programmable bench power supply is designed with taking into consideration the following features and requirements:
Fig. 1: EEZ PSU early prototype
- Modular design that allows combining modules with various performance and capability and creation of multiple output solution (min. two channels)
- Single input for both power and bias supply removes needs for custom multiple secondary winding main transformer
- SMPS pre-regulator circuit for lower power dissipation (using e.g. SMPS)
- SMPS pre-regulator's "100% Duty cycle" mode of operation to lower output ripple and noise
- TFT color touch-screen as input device and informative data output
- MCU managed serial and parallel connection if two channels channels with same capability exists
- Voltage regulation (CV), 1/10 mV resolution, up to 50 V (default: 0 – 40 V)
- Current regulation (CC), 0.1/1 mA resolution (when dual range, default: 0 – 500 mA / 0 – 5 A)
- Minimal overshoot when new voltage or current level is programmed or in marginal cases such as power on/off
- Min. 15-bit data acquisition resolution
- Various protection mechanisms: over-voltage (OVP), over-current (OCP), over-power (OPP) and over-temperature (OTP), current limiting when temperature sensors or cooling fan fault is detected
- Switching frequency synchronization between switching regulators
- Output enable (OE) circuit
- Remote voltage sensing
- Remove voltage programming (“analog” programming)
- External digital trigger (3.3 or 5 V logic level)
- Simple AC input protection (surge and transient protection)
- Selectable 115/230 VAC, Soft-start/Stand-by mode
- Simple DC output protection (reverse voltage, over-voltage)
- Serial and parallel firmware controlled connection of two power outputs
- MCU based digital control module using 32-bit Arduino/Genuino boards such as Due (or possibly Star - OTTO in the future)
- SPI bus for communication with peripherals
- Galvanic isolation between digital control module and power channels
- Real-time clock (RTC) with supercap/batter backup
- EEPROM used as a configuration and calibration parameters storage
- SD-card as an additional storage
- USB support that can be used for remote control, as a debug console and firmware upload
- Ethernet support for remote control
- Fan speed control for “quiet” operation
- Incremental encoder with tactile switch as an additional mean of setting output values
- Usage of easy to find components that could be obtain with just single order from suppliers such as Farnell element14, TMU.eu, DigiKey, Mouser, RS online, etc.
- Possibility to house at least dual channel version in affordable and compact (2U height) metallic enclosure
- Comprehensive software support that include SCPI remote control, GUI configuration and monitoring and even cross platform (Windows, Linux, OS X) software simulator of the firmware functionality
Despite the fact that proposed design is rather complex it should not be a problem to follow it and build in few steps at home with intermediate soldering skills of SMT components. Required layouts (Eagle and Gerbers) for PCB production are available on GitHub or PCBs can be directly ordered e.g. from OSHpark.
Firmware uploading requires novice programming skills and familiarity with Arduino IDE (check also Videos section for examples).
## The PSU building blocks and default configuration
The power supply’s modular design gives us more freedom in building functional device. The following modules is possible to use for PSU building:
- Auxiliary power module,
- Power board (one per channel) and
- Arduino Shield (digital control)
Each module will be described in more details in text that follows. The default configuration is named EEZ H24005 and comes with the following features:
- Two “floating” channel with programmable output voltage range 0 – 40 V and current 0 – 5 A (max. 155 W per channel) or 0 – 80 V and 0 – 10 A with MCU managed connection in series or parallel
- Remote voltage sensing (with inverse polarity protection), remote voltage programming
- External digital trigger input, external temperature sensor
- Temperature controlled cooling fan
- AC mains power comes through two AC/DC power module (48 Vdc, 155 W)
- Arduino Due based digital control
- 3.2” TFT color touch-screen as local console
- Ethernet and USB for remote control (USB is used for firmware update)
- Metal enclosure: 80 (H) x 290 (W) x 240 (D without handles) mm
## Functional block diagram
The block diagram of a PSU is shown on the picture below. It will require four modules in total. The auxiliary power module used for autonomous powering of the Arduino Shield. Complete control of all functionality is performed using local display with touch screen as an input device. In addition the Arduino shield enables that fully featured remote control is established using serial (via USB) or Ethernet communication.
The remote control is based on a command set that is written in compliance with the SCPI 1999.0 specification that is de facto standard for the middle and high-end commercial power supplies and many other laboratory instruments.

## Module description
Fig. 2: EEZ PSU functional block diagram
More information about modules used for building the PSU can be found in chapters that follows. A header section with information about module development status was added to simplify progress tracking, for example:

Current version	r4B39
Status	Completed, ready for production
PCB manufactured	Yes
PCB assembled	Yes
BOM	Yes (Farnell, TME)
File repository	https://github.com/eez-open/psu-hw/tree/master/Aux%20power