# EEZ PSU H24005 wire harness
- Ethernet cable
- Cooling fan cable
- AUX temperature sensor cable
- AC switch patch cables
- Cable set #1
- Cable set #2
- Pre-assembled cables
- BOM

## Ethernet cable
Total length	260 (± 5) mm
Type	4 pairs Ethernet, stranded wire
TME	HELUKAT100-FTP-FP or e.g. cut CA00600.00C01 to size
Farnell	1734894 or e.g. cut 1602385 to size
Connectors
Side A	X8	Side B	X18
Housing type	Receptacle, 0.1”	Housing type	See side A
TME	NSR-08	TME	See side A
Farnell	865667	Farnell	See side A
Pin:	Female	Pin	See side A
Pin quantity	8	Pin quantity	7
TME	NDR-T	TME	See side A
Farnell	1022219	Farnell	See side A

Wiring
Name	Position	Color / description	Position	Color / description
LED link	1	Brown-white	1	Brown-white
LED act	2	Brown	2	Brown
Shield	3	Blue-white	3	NOT CONNECTED
RX+	4	Orange-white	4	Orange-white
RX-	5	Orange	5	Orange
TX+	6	Green-white	6	Green-white
TX-	7	Green	7	Green
+3.3V	8	Blue	8	Blue

Picture

## Cooling fan cable
Total length	100 (± 5) mm
Type	3 wire, already mounted on the cooling fan
Fan:	TME: MB60251V1-G99
Connectors
Side A	X4	Side B	N/A
Housing type	Receptacle, 0.1”	Housing type	N/A
TME	MX-6471-031	TME	N/A
Farnell	143127	Farnell	N/A
Pin	Female	Pin	N/A
Pin quantity	3	Pin quantity	0
TME	MX-0850-0114	TME	N/A
Farnell	2063734	Farnell	N/A
Additional info
Fan is already equipped with cable that has to be cut to size and terminated with connector

Wiring
Name	Position	Color / description
12 Vdc	1	Red
GND	2	Black
Sense	3	Yellow

Picture

## AUX temperature sensor cable
Total length	120 (± 5) mm
Type	2 wire, already mounted on the temperature sensor
NTC	TME: NTCLE100E3103JB0
Wire	TME: 9921.004100
Connectors
Side A	X17	Side B	N/A
Housing type	Receptacle 0.1”	Housing type	N/A
TME	MX-6471-021	TME	N/A
Farnell	143126	Farnell	N/A
Pin	Female	Pin	N/A
Pin quantity	2	Pin quantity	0
TME	MX-0850-0114	TME	N/A
Farnell	2063734	Farnell	N/A
Additional info
The temperature sensor is already equipped with cable that has to be terminated with connector

Wiring
Name	Position	Color / description
MCU_VCC	1	Yellow
NTC	2	Yellow

Picture

## AC switch patch cables
Total length	45 (± 5) mm
Type	1.5 mm2 singe wire, stranded, min. 300 VAC
TME	Brown: LGY1.5/10-BR, Blue: LGY1.5/10-BL
Farnell	Brown: 1764967, Blue: 1764966
Connectors
Side A	AC output	Side B	AC Switch input
Pin	Female quick disconnect	Pin	See side A
Pin quantity	1	Pin quantity	See side A
TME	ST-003/05-20T	TME	See side A
Farnell	2060332	Farnell	See side A
Additional info
Required for Schurter 4304.6090 IEC inlet (C14, 6765, 10A)

Wiring1
Name	Position	Color / description	Position	Color / description
Line	1	Brown	1	Brown

Wiring2
Name	Position	Color / description	Position	Color / description
Neutral	1	Blue	1	Blue

Picture
N/A
## Cable set #1

### Section A: Power board X1 cable
Length	400 (± 5) mm
Type	1.5 mm2 singe wire, stranded, min. 300 VAC
TME	Red: LGY1.5/10-RD, Black: LGY1.5/10-BK / Green-Yellow: LGY1.5/10-YL/GR
Farnell	Red: 1764971, Black: 1764965, Green-Yellow: 1764969
Connectors
Side A	Power board X1 (Ch2)	Side B	AC/DC module power output
Housing type	Male, Mini-Fit Jr.	Housing type	Ferrule, single wire
TME	MX-5557-03R2	TME	H1.5/14D/BK
Farnell	9963480	Farnell	9972307
Pin	Female	Quantity	2
Pin quantity	3	Housing type	Ferrule, twin entry
TME	MX-5556-T3L	TME	H1.5/16D2/BK
Farnell	9732675	Farnell	1121754
		Quantity	1

Wiring
Name	Position	Color / description	Position	Color / description
Vin+	1	Red	N/A	Ferrule
Vin-	2	Black	N/A	Ferrule
Earth	3	Green-Yellow	N/A	Ferrule, twin entry

### Section B: AC/DC module Earth
Length	70 (± 1) mm
Type	1.5 mm2 singe wire, stranded, min. 300 VAC
TME	LGY1.5/10-YL/GR
Farnell	1764969
Connectors
Side A	AC/DC module Earth output	Side B	AC/DC module Earth (in)
Housing type	Use twin entry ferrule, See B side, Section A	Housing type	Ferrule, twin entry
		TME	H1.5/16D2/BK
		Farnell	1121754
		Quantity	1

### Section C: Chassis Earth
Length:	170 (± 5) mm
Type:	1.5 mm2 singe wire, stranded, min. 300 VAC
TME	LGY1.5/10-YL/GR
Farnell	1764969
Connectors
Side A	AC/DC module Earth (in)	Side B	Chassis Earth (bottom plate)
Housing type	Use twin entry ferrule, See B side, Section B	Housing type	Non-insulated, Ring terminal, 4 mm
		TME	3804/M.05.00.9
		Farnell	-
		Quantity	1

### Section D: Arduino Shield Earth
Length:	500 (± 5) mm
Type:	1.5 mm2 singe wire, stranded, min. 300 VAC
TME	LGY1.5/10-YL/GR
Farnell	1764969
Connectors
Side A	Chassis Earth (bottom plate)	Side B	Arduino Shield Earth tab (J6)
Housing type	Use ring terminal, See B side, Section C	Housing type	Female quick disconnect
		TME	ST-003/05-20T
		Farnell	2060332
		Quantity	1

## Cable set #2

### Section A: Power board X1 cable
Length	330 (± 5) mm
Type	1.5 mm2 singe wire, stranded, min. 300 VAC
TME	Red: LGY1.5/10-RD, Black: LGY1.5/10-BK / Green-Yellow: LGY1.5/10-YL/GR
Farnell	Red: 1764971, Black: 1764965, Green-Yellow: 1764969
Connectors
Side A	Power board X1 (Ch1)	Side B	AC/DC module power output
Housing type	Male, Mini-Fit Jr.	Housing type	Ferrule, single wire
TME	MX-5557-03R2	TME	H1.5/14D/BK
Farnell	9963480	Farnell	9972307
Pin	Female	Quantity	2
Pin quantity	3	Housing type	Ferrule, twin entry
TME	MX-5556-T3L	TME	H1.5/16D2/BK
Farnell	9732675	Farnell	1121754
		Quantity	1

Wiring:
Name	Position	Color / description	Position	Color / description
Vin+	1	Red	N/A	Ferrule
Vin-	2	Black	N/A	Ferrule
Earth	3	Green-Yellow	N/A	Ferrule, twin entry

### Section B: AC/DC module Earth
Length	150 (± 5) mm
Type	1.5 mm2 singe wire, stranded, min. 300 VAC
TME	LGY1.5/10-YL/GR
Farnell	1764969
Connectors
Side A	AC/DC power module Earth (in)	Side B	AC inlet Earth
Housing type	Use twin entry ferrule, See B side from Section A	Housing type	Female quick disconnect
		Quantity	1
		TME	ST-003/05-20T
		Farnell	2060332

### Section C: AUX PS module Earth
Length	150 (± 5) mm
Type	1.5 mm2 singe wire, stranded, min. 300 VAC
TME	LGY1.5/10-YL/GR
Farnell	1764969
Connectors
Side A	AC inlet Earth	Side B	AUX PS module (X4)
Housing type	Use Female quick disconnect terminal, See B side, Section B	Housing type	Male, Mini-Fit Jr.
		TME	MX-5557-05R2
		Farnell	9963502
		Pin	Female
		Pin quantity	1
		TME	MX-5556-T3L
		Farnell	9732675

Wiring:
Name	Position	Color / description	Position	Color / description
-	-	Female quick disconnect	5	Earth

### Section D: Lower AC/DC module power input
Length	200 (± 5) mm
Type	1.5 mm2 singe wire, stranded, min. 300 VAC
TME	Brown: LGY1.5/10-BR, Blue: LGY1.5/10-BL
Farnell	Brown: 1764967, Blue: 1764966
Connectors
Side A	AUX PS module (X3)	Side B	Lower AC/DC module power input
Housing type	See Side B, Section C	Housing type	Ferrule, twin entry
TME	-	TME	H1.5/16D2/BK
Farnell	-	Farnell	1121754
Pin	Female	Quantity	2
Pin quantity	2
TME	MX-5556-T3L
Farnell	9732675

Wiring:
Name	Position	Color / description	Position	Color / description
Line	1	Brown	N/A	Ferrule, twin entry
Neutral	2	Blue	N/A	Ferrule, twin entry

### Section E: Upper AC/DC module power input
Length	70 (± 1) mm
Type	1.5 mm2 singe wire, stranded, min. 300 VAC
TME	Brown: LGY1.5/10-BR, Blue: LGY1.5/10-BL
Farnell	Brown: 1764967, Blue: 1764966
Connectors
Side A	Lower AC/DC module power input	Side B	Upper AC/DC module power input
Housing type	Use twin entry ferrule, See B side, Section D	Housing type	Ferrule, single wire
		TME	H1.5/14D/BK
		Farnell	1905292
		Quantity	2

### Section F: AUX PS module power input
Length	110 (± 5) mm
Type	1.5 mm2 singe wire, stranded, min. 300 VAC
TME	Brown: LGY1.5/10-BR, Blue: LGY1.5/10-BL
Farnell	Brown: 1764967, Blue: 1764966
Connectors
Side A	AUX PS module (X3)	Side B	Lower AC/DC module power input
Housing type	See Side B, Section C	Housing type	Female quick disconnect
TME	-	TME	ST-003/05-20T
Farnell	-	Farnell	2060332
Pin	Female	Quantity	2
Pin quantity	2
TME	MX-5556-T3L
Farnell	9732675

Wiring:
Name	Position	Color / description	Position	Color / description
Line	3	Brown	N/A	Female quick disconnect
Neutral	4	Blue	N/A	Female quick disconnect

## Pre-assembled cables
### Arduino shield power cable

Length	300 mm
Type	Ribbon cable with IDC connectors; 10x28AWG
TME	FC10300-0
Farnell	N/A

Picture

### Arduino USB cable
Length	300 or 500 mm
Type	USB Micro to 5-pin receptacle cable
TME	N/A
Farnell	N/A
AliExpress	e.g. High Quality 50CM micro usb 5pin male to 1x 5Pin Female 0.1" USB header PCB motherboard cable

## BOM
Type	Name	Total
Receptacle	NSR-08	2
Receptacle	MX-6471-031	1
Receptacle	MX-6471-021	1
Receptacle Mini-Fit Jr.	MX-5557-03R2	2
Receptacle Mini-Fit Jr.	MX-5557-05R2	1
Pin	NDR-T	15
Pin	MX-0850-0114	5
Pin	MX-5556-T3L	11
Female quick disconnect	0	8
Ferrule	H1.5/14D/BK	6
Ferrule, twin entry	H1.5/16D2/BK	5
Ring terminal	3804/M.05.00.9	1
Cat5e cable	HELUKAT100-FTP-FP	0.26 m
1.5 cm2 cable (brown)	LGY1.5/10-BR	0.47 m
1.5 cm2 cable (blue)	LGY1.5/10-BL	0.47 m
1.5 cm2 cable (red)	LGY1.5/10-RD	0.73 m
1.5 cm2 cable (black)	LGY1.5/10-BK	0.73 m
1.5 cm2 cable (yellow-green)	LGY1.5/10-YL/GR	1.77 m
0.35 cm2 cable	9921.004100	0.24 m
12Vdc cooling fan	MB60251V1-G99	1
NTC	NTCLE100E3103JB0	1
IDC ribbon cable	FC10300-0	0.3 m
USB Micro cable	Search Aliexpress	0.5 m